## Sheet5
yet undefined | [16 empty] | Beetles count from 1-3, Ports count 1-4
module seems working o.k (noise, Temp sensor)
successful commissioned with testpulse | [14 empty] | IT1 A | X2(3) | Beetle2 port 1,3 are garbage (testpules have about ½ of hight) --> tell1 problem/ittell9 -->Fixed 24Jun08
error... see description
a minor anomaly but probably fine | [14 empty] | IT1 A | U(5) | Beetle3 Port1 slightly larger/structured noise (2.8ADC Digiboard??)
some mapping problem (fibres on Digiboard) | [14 empty] | IT1 A | U(4) | overall slightly (0.2ADC) and NON FLAT noise
IT1 Top Box
|  |  |  | X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7) |  |  |  |  | IT1 B | U(3) | Beetle3 Port1 dead (seen only 28May not 26May, not 1xJune...
|  |  |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  |  |  |  | IT1 B | U(4) | Beetle1 Port3 slightly larger/structured noise (2.2ADC Digiboard?? --> now 6ADC in June)
|  |  |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) |  |  |  |  | IT1 B | U(5) | Beetle3 Port2 slightly larger/structured noise (2.2ADC Digiboard??)
|  |  |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7)
[15 empty] | IT1 C | X1(2) | Beetle1 Port2 ZERO noise (no testpulse) --> hybrid cable swapped-->problem stays--> Digiboard downwards
|  | IT1 Cryo Box |  | U(7) V(7): IT1 Access Box | [6 empty] | ittell04/PP0/link3 (parts of it.. NO problem in TELL1 BER test, guess it is the digiboard
X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7) |  | X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7)
U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) | IT1 C | U(1) | Beetle2 Port4 low noise (no testpulse) re-plugged --> now the tell1 is off, cannot tell yet --> ittell3
V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) | IT1 C | V(6) | Problems with configuring the beetles, first appeared 12June
Same module also has problems reading the hybrid temp. (reads some 200C :))
X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7) |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7) | IT1C | X2(7) | Problems with configuring the beetles, first appeared 17June --> ittell1
[15 empty] | IT1 T | V(6) | Beetle3 Port3 low noise (no testpulse) re-plugged--> still the same --> ittell7
U(7) V(7): IT1 Bottom Box | [8 empty] | IT1 T | V(6) | Beelte3 Port4 has 10% larger gain for testpulse ...
U(5) V(5): X1(1); U(6) V(6): X1(2); U(7) V(7): X1(3) | X1(4) | X1(5) | X1(6) | X1(7) |  |  |  |  | IT1 T | V(4) | Beetle2 Port2 low noise (no testpulse)-->ittell6 --> couldn't be seen anymore after Tell1 exchange (24Jun08)
|  |  |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  |  |  |  | IT1 T | V(3) | Beetle3 Port2 slightly larger and structured noise (2-2.4ADC, NOT DIGIBOARD NOISE, moved when swapping cables)
|  |  |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7)
|  |  |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7)
Problematic optical links correspond to:
ittell9 PP3 link1 == A X2(3)
ittell10 PP0 link3 == A U(6) Beetle 1
ittell14 PP3/Link0 == B X2(2) Beetle 1
ittell5 PP3 link2 (traced to tell1 input) == T U(2) Beetle3
Kink in fibre A-side SvcBox4, Board15, Beetle2 --> A X1(7), ittell11 PP1/Link4
Kink in fibre A-side SvcBox4, Board15, Beetle2 --> B U(7), ittell12 PP3/Link5
IT2 Top Box
|  |  |  | X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7)
|  |  |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7)
|  |  |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) |  |  |  |  | IT2 T | V(1-4) X2(1-4) | Connected to ittell20 with stuck even id bit--> excluded.. (swapped with ittell19 to check the corresponding modules)
|  |  |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7)
|  | IT2 Cryo Box |  | U(7) V(7): IT2 Access Box |  |  |  |  | IT2 T | U(6) | Beetle1 Port2 low noise/low testpulse (10ADC)-->ittell21
X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7) |  | X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7) | IT2 T | X2(6) | Beetle1 Port4 low noise/no testpulse-->ittell21
U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7)
V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7)
X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7) |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7) | IT2 T | X2(4) | Beetle1 Port4 low noise, LOW testpulse--> checked and re-plugged digiboard connector. Problem remained (13June) --> ittell20
[15 empty] | IT2 C | V(6) | Beetle1 Port4 HUGE noise (80-100 w/o CMS, 30-80 CMS)-->swapped with V(5)/problem stayed in V(6)-->Problem lies Digiboard downwards
/No testpulse (nothing special on ittell16,PP1/Link0 in BER test)
IT2 Bottom Box
U(5) V(5): X1(1); U(6) V(6): X1(2); U(7) V(7): X1(3) | X1(4) | X1(5) | X1(6) | X1(7) |  |  |  |  | IT2 C | U(7) | strange larger/structured noise (2.5-3 ADC counts)-->swapped with U(6)/problem stayed in U(7)-->Problem lies in Digiboard downwards (13June)
|  |  |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  |  |  |  | IT2 C | X2(4) | Beetle2, Port4, low noise(0.9ADC) low testpulse(15ADC) new since 10June. -->ittell16
|  |  |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) |  |  |  |  | IT2C | X2(3) | temp sensor reading of -42deg (19.June)...
|  |  |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7)
[15 empty] | IT2 A | X2(5) | Beetle2 Port2 large/structured noise (Digiboard?)
[15 empty] | IT2 A | V(4) | Beetle3 Port1 large/structured noise (Digiboard?)
[15 empty] | IT2 A | X1(1) | probably swapped fibres on Digiboard --> fixed (12June, fibre in cable chain internally crossed--> replaced by spare)
[15 empty] | IT2 A | X1(2) | probably swapped fibres on Digiboard --> fixed (12June, fibre in cable chain internally crossed--> replaced by spare)
[15 empty] | IT2 A | X1(3) | probably swapped fibres on Digiboard --> fixed (12June, fibre in cable chain internally crossed--> replaced by spare)
[15 empty] | IT2 A | U(4) | Beetle3 (-->ittell22, problems with optical link, traced to within the SvcBox)
[15 empty] | IT2 A | X1(7) | Beetle1,2 (-->ittell25, problems with optical link, traced to within the SvcBox)
IT3 Top Box
|  |  |  | X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7)
|  |  |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7)
|  |  |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7)
|  |  |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7)
[15 empty] | IT3 C | U(3) | Beetle3, port4, large noise peak (5ADC)in first 15channels
--> DIGIBOARD NOISE
|  | IT3 Cryo Box |  | U(7) V(7): IT3 Access Box
X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7) |  | X1(1) | X1(2) | X1(3) | X1(4) | X1(5) | X1(6) | X1(7) | IT3 C | U(4) | very large common mode noise (50ADC counts)
U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  |  | --> MODULE (L229/SideBoxC1/A5_4/showed CMN of up to 7ADC only..)
V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7)
X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7) |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7) | IT3 C | U(6) | temp sensor always above 42 deg.
Cable swapping at SvcBox showed problem within Module (detector box)
U(7) V(7): IT3 Bottom Box | [8 empty] | IT3 T | X1(7) | FE does not configure (connects to ittell32)-->FIXED (re-plugging cable), BUT since 12June, after a while the analog current goes up and sometimes goes into error
U(5) V(5): X1(1); U(6) V(6): X1(2); U(7) V(7): X1(3) | X1(4) | X1(5) | X1(6) | X1(7)
|  |  |  | U(1) | U(2) | U(3) | U(4) | U(5) | U(6) | U(7) |  |  |  |  | IT3 T | V(6) | FE configures only some days (connects to ittell35)
|  |  |  | V(1) | V(2) | V(3) | V(4) | V(5) | V(6) | V(7) | [6 empty] | --> problem traced to Digiboard/SvcBox NOT the MODULE/Cable, module tested to be o.k.
|  |  |  | X2(1) | X2(2) | X2(3) | X2(4) | X2(5) | X2(6) | X2(7)
[15 empty] | IT3 T | U(1) | last two bits in PCN missing (from 4th port of first beetle, visible in PCN and noise)-
=-> this port also does not see any testpulse! -->digiboard-->ittell33
[15 empty] | IT3 T | V(1) | 3rd beetle 1st port larger noise (2.7ADC with structure),
in the testpulse you see a large spike in the middle of this port --> DIGIBOARD NOISE
[15 empty] | IT3 B | X2(7) | 3rd beetle no data --> corresponds to ittell41/PP3 link 5-->FIXED
[15 empty] | IT3 B | V(6) | 2nd and 3rd beetle no data --> corresp. ittell41/PP1/link1,2--> FIXED
[15 empty] | IT3 B | U(3) | 3rd beetle 3rd port slightly lower noise -> testpulse arrives
[15 empty] | IT3 A | X1(5) | 3rd beetle 4th port LOW noise --> NO testpulse --> ittell39
[15 empty] | IT3 A | X1(5) | 29May: Whole module has low testpulse (22 instead 40in RMS)

## Tests St1
[7 empty] | Control tests |  |  |  |  | HV tests |  | Test pulse (Vfs=400, I=0x2c)
| General comment | Tell1 | Ladder | RO sector | Box | Dig Board | Ctrl Board | Dig Board | Ladder | ORX test | Tell 1 | I meas. | CMS noise | max | +25ns | remainder | Remarks
ST1A | 28.3 Frequent SPECs communication problems with.. (seen when programing DATEPROM from the SPECS Client pannel) 31.3.SPECS communication is fine if we make the chaining 2 3 4 1 instead of 1 2 3 4. Probably the SPECS bus is on the limit concerning the cable length and the resulting signal integrity, and the receiver in SvcBox2 is slightly more tolerant. At the end of the chain, SvcBox1 gets a "refreshed" signal. Note: Detector boxes are not in place, henc e no proper grounding of Control board. | ittell08 | [7 empty] | OK
|  | ittell09 | [7 empty] | FAIL | [6 empty] | PP3->link1 Bit error:0x000000 Missing word:0xFF -> The problem has been traced up to the TELL1 input.
|  | ittell10 | [7 empty] | FAIL | [6 empty] | PP0->link3 error0xFFFFFF: completely bad -> Problem traced to be in SB4 DB2 GOL1 link inside Svce box
|  | ittell11 | [7 empty] | OK
|  | ittell12 | [7 empty] | OK
|  | ittell13 | [7 empty] | OK
|  | ittell14 | [7 empty] | FAIL | FAIL | [5 empty] | ittell14 has an ID of 279, the GBE card on this tell1 board has an I2C error. 4 bit errors in PP3/Link0 observed (in about 10min)
ST1C | 28.3. SPECS communication is fine | ittell01 |  |  | BOX CRYO |  |  |  |  | OK
|  | ittell02 | [7 empty] | OK
|  |  | s6v |  |  | sb1sl3 |  |  | Temperature read Error
|  | ittell03 | [7 empty] | OK | FAIL | [5 empty] | some strange problem, error message when trying to take data. When swapping fibre input, it remaines at the same TELL1. Hence should be caused by the tell1, not the detector data
|  | ittell04 | [7 empty] | OK
BOX TOP
|  | ittell05 | [7 empty] | FAIL | [6 empty] | PP3->link2 error0xFFFFFF: completely bad -> The problem has been traced up to the TELL1 input.
|  | ittell06 | [7 empty] | OK | [6 empty] | ittell06 has a bit stuck at "1" in the Dest IP address
|  | ittell07 | [7 empty] | OK

## Tests St2
[7 empty] | Control tests |  |  |  |  | HV tests |  | Test pulse (Vfs=400, I=0x2c)
| General comment | Tell1 | Ladder | RO sector | Box | Dig Board | Ctrl Board | Dig Board | Ladder | ORX test | Tell 1 | I meas. | CMS noise | max | +25ns | remainder | Remarks
ST2A | 28.3. SPECS communication is fine | ittell22 | [7 empty] | Fail | OK | [5 empty] | PP3->link5 ~500err/10s: problem remained after cleaning 17.3. (12June): swapped cables in bunker and svcbox-->problem traced to INSIDE SvcBox. It was clear, as it even allowed us to pin down the crossed fibres in the cable chain connecting to ittel22UP
|  | ittell23 | [7 empty] | Fail | OK | [5 empty] | PP0->link4 ~16err/5min: problem remained after cleaning (maybe slightly better, 2 errors/5mnin) (12June: problem could not be seen anymore after swapping up/down)
|  | ittell24 | [7 empty] | OK | OK
|  | ittell25 | [7 empty] | Fail | OK | [5 empty] | PP1->link3 completely bad PP1->link4 ~600 err/1min both problems remained after cleaning 17.3./ when swapping up/down links in the bunker (12June), the problem (Link3) swapped --> located cablechain upwards. (link4 could not be seen anymore). Then exchanging the fibre in the cable chain (both ends) the problem stayed where it is. --> located inside SvcBox
|  | ittell26 | [7 empty] | OK | OK
|  | ittell27 | [7 empty] | OK | OK
|  | ittell28 | [7 empty] | Fail | OK | [5 empty] | PP3 link 0,1,2 now in error (18.3.)… tried re-programming and then even more links fail…. Problem of programming the GOLs ??.. get's even worse (30May), now PP0Link3 and PP3/Link1,2 are in error, the problem was funny. It seems to be a GOL programming thing, as with different intialisations, sometimes only link1, 1+2, 1+2+3 show the problem. HOWEVER, swapping the fibres (in bunker, up<->down, we even veryfied that we swap the correct fibres) the link problem STILL remained in the same TELL1 PP3 and didn't move to PP1 (12June)
ST2C | The svce boxes were overheating, the mixed water handle was closed. Servicebox 4, Partition1 (1-4), I2C line for programming the beetle doesn't work. Servicebox 2 shows startup current consumption problem. Inhibit pull up resisitor that switches of fhe LV regulators seems to be open. Ok after Control Card initialisation as then this takes care of the switchin on/off. 28.3. Added "pull up cable" to fix I2C (SPECS) communicatio problem... then SPECS communication worked for all 4 boxes, but.. 28.3. Service boxes (SPECS communication) stopped working.. on and off.. then completely... 4.4.After testing the long and short distance cables (and spares) which all seem to have very much the same characteristic, the problem could not be reproduced anymore... | ittell15 | [7 empty] | OK | OK
|  | ittell16 | [7 empty] | OK | OK
|  | ittell17 | [7 empty] | OK | OK
|  | ittell18 | [7 empty] | OK | OK
|  | ittell19 | [7 empty] | OK | OK | [5 empty] | ittell19 reports an error when programming it. Fixed, tell1 exchanged
|  | ittell20 | [7 empty] | OK | OK | [5 empty] | ittell20 didn't pass the selftest. (could not be reproduced..seems fine) The bottom half opt links are all bad: LOS+TLK status ERROR. Fixed (16.4.2008.. Cable was not properly plugged in bunker). 9.May: ittell20 escaped the proper programing of the recipies (although I didn't see an error reported) and it seems to have a stuck bit in the event number that it sends. SOMETIMES the bit corresponding to 0x40 is stuck in the event number.... (Olivier Callot)
|  | ittell21 | [7 empty] | OK | OK

## Tests St3
[7 empty] | Control tests |  |  |  |  | HV tests |  | Test pulse (Vfs=400, I=0x2c)
| General comment | Tell1 | Ladder | RO sector | Box | Dig Board | Ctrl Board | Dig Board | Ladder | ORX test | Tell 1 | I measured | CMS noise | max | +25ns | remainder | Remarks
ST3A | SPECS communication doesn't work with Svce box 4, it works for the first 3 -> need to check the SPECS link cable and/or ctrl board. -> Fixed, Ctrl Board SPECS cable was connected to the output instead of input. 28.3. SPECS communication is fine 28.3 SvcBox 4 voltage reference extremely low (1.9V should be 2.5V) Possible reset problem in the DCU. 31.3.Optical Links (ittell41 u/d) swapped with ittell40 (u/d) on bunker patch pannel --> problem swapped --> problem located between bunker and station. | ittell36 | [7 empty] | OK
|  | ittell37 | [7 empty] | OK
|  | ittell38 | [7 empty] | OK
|  | ittell39 | [7 empty] | OK
|  | ittell40 | [7 empty] | Ok
|  | ittell41 | [7 empty] | FAIL | [6 empty] | PP1 -> link1,2 completely bad (LOS)Fixed(spare) 27.May08 PP2 -> link2 100 err/min. (27May08: 1 error in 30min .. why better? PP3 -> link5 completely bad (LOS) Fixed(spare) 27.May08
|  | ittell42 | [7 empty] | OK
ST3C | SPECS communication doesn' t work with Svce boxes 3 and 4, it works for 1 and 2 -> need to check the SPECS link cable and/or ctrl board. --> (28.3) jumper for terminating the SPECS chain were in the "terminatating" position for SvcBox 2 (and 4) and is now put in place correctly The svce boxes were overheating, the mixed water handle was closed. 28.3. SPECS communication is fine i | ittell29 | [7 empty] | OK
|  | ittell30 | [7 empty] | OK
|  | ittell31 | [7 empty] | OK
|  | ittell32 | [7 empty] | OK | [6 empty] | LX1(7) could not be configured (4.4.)
|  | ittell33 | [7 empty] | OK
|  | ittell34 | [7 empty] | OK
|  | ittell35 | [7 empty] | OK

## Tell1 status
Tell1 | Status
1
2
3 | MEPRx001: Bad packet from > ittell03 [mismatch between length on wire and in MEP header] MEPRx001: Bad packet from > ittell03 [mismatch in packing factor]--> exchanged 24June06
4
5 | ORX input error: PP3->link2 error0xFFFFFF: completely bad -> The problem has been traced up to the TELL1 input.--> exchanged 24June06
6 | Destination IP problem: one of the bits is stuck at “1”, for ex, instead of showing 200.1, it shows 232.1. Gong is looking at it.--> exchanged 24June06
7
8
9 | Problems with input from the ORX card--> exchanged 24June06
10
11
12
13
14
15
16
17
18
19 | Error when trying to reset/program the GBE. Gong is looking at it. --> fixed
20 | stuck bit in event ID--> exchanged 24June06
21
22
23
24
25
26
27
28 | Traced to TELL1 in put as failure (need to repeat swapping at tell1 input) but link failure (PP3 Link0,1,2) does not allways happen. e.g. after re-configuring the Frontend (GOL) sometimes some links don't show the problem anymore. 18June: if input is swapped with ittell27, none of the two show any error (repeated and checked 3 times)--> input seem to have different sensitivity and transmitter slightly lower power. --> exchanged 24June06
29
30
31
32
33
34
35
36
37
38
39
40
41
42
Note | exchanged TELL1 boards 24.June 2008
ittell03 new 00:30:59:01:C3:E5 Tell1_id=352
ittell05 new 00:30:59:01:C7:E8 Tell1_id=060
ittell06 new 00:30:59:01:C7:B6 Tell1_id=003
ittell09 new 00:30:59:01:C4:3E Tell1_id=014
ittell20 new 00:30:59:01:C8:AD Tell1_id=210
ittell28 new 00:30:59:01:C8:A7 Tell1_id=159

## Listing
Production Module number | Last digits of the barcode | Box position | Layer | sector | strip numbers of shorts (separated by comma) | strip numbers of open ch. (separated by comma) | strip numbers of pinholes (separated by comma) | Tell1 SubID | Tell1 SourceID | Tell1 link | Tell1 Name | Module Postion name (local to box)
L188 | 186 | IT1Access | U | 1 |  |  |  | 7 | 7 | 4 | ittell8 | B5(7)
L108 | 187 | IT1Access | U | 2 |  |  |  | 7 | 7 | 5 | ittell8 | B5(6)
L185 | 188 | IT1Access | U | 3 |  |  |  | 7 | 7 | 6 | ittell8 | B5(5)
L184 | 189 | IT1Access | U | 4 |  |  | 381 | 7 | 7 | 7 | ittell8 | B5(4)
L113 | 190 | IT1Access | U | 5 |  |  | 351 | 10 | 10 | 1 | ittell11 | B5(3)
L112 | 191 | IT1Access | U | 6 | [156/157];[160/161] |  |  | 9 | 9 | 1 | ittell10 | B5(2)
L197 | 192 | IT1Access | U | 7 |  | 6 |  | 9 | 9 | 3 | ittell10 | B5(1)
L196 | 171 | IT1Access | V | 1 |  |  |  | 8 | 8 | 0 | ittell9 | A5(7)
L193 | 170 | IT1Access | V | 2 |  |  |  | 8 | 8 | 1 | ittell9 | A5(6)
L192 | 169 | IT1Access | V | 3 |  |  |  | 8 | 8 | 2 | ittell9 | A5(5)
L189 | 168 | IT1Access | V | 4 |  |  |  | 8 | 8 | 3 | ittell9 | A5(4)
L117 | 167 | IT1Access | V | 5 |  |  |  | 9 | 9 | 0 | ittell10 | A5(3)
L116 | 166 | IT1Access | V | 6 |  |  |  | 9 | 9 | 2 | ittell10 | A5(2)
L109 | 165 | IT1Access | V | 7 |  |  |  | 9 | 9 | 6 | ittell10 | A5(1)
L183 | 179 | IT1Access | X1 | 1 |  |  |  | 7 | 7 | 0 | ittell8 | B0(7)
L182 | 180 | IT1Access | X1 | 2 |  |  |  | 7 | 7 | 1 | ittell8 | B0(6)
L111 | 181 | IT1Access | X1 | 3 |  |  |  | 7 | 7 | 2 | ittell8 | B0(5)
L119 | 182 | IT1Access | X1 | 4 |  |  |  | 7 | 7 | 3 | ittell8 | B0(4)
L194 | 183 | IT1Access | X1 | 5 |  |  |  | 10 | 10 | 0 | ittell11 | B0(3)
L107 | 184 | IT1Access | X1 | 6 |  |  |  | 10 | 10 | 2 | ittell11 | B0(2)
L186 | 185 | IT1Access | X1 | 7 |  |  |  | 10 | 10 | 3 | ittell11 | B0(1)
L110 | 178 | IT1Access | X2 | 1 |  |  |  | 8 | 8 | 4 | ittell9 | A0(7)
L195 | 177 | IT1Access | X2 | 2 |  |  |  | 8 | 8 | 5 | ittell9 | A0(6)
L191 | 176 | IT1Access | X2 | 3 |  |  |  | 8 | 8 | 6 | ittell9 | A0(5)
L190 | 173 | IT1Access | X2 | 4 |  |  |  | 8 | 8 | 7 | ittell9 | A0(4)
L118 | 174 | IT1Access | X2 | 5 | [2/3] |  |  | 9 | 9 | 4 | ittell10 | A0(3)
L115 | 173 | IT1Access | X2 | 6 |  |  |  | 9 | 9 | 5 | ittell10 | A0(2)
L114 | 172 | IT1Access | X2 | 7 |  |  |  | 9 | 9 | 7 | ittell10 | A0(1)
L73 | 7 | IT1Bottom | U | 1 |  |  |  | 13 | 13 | 0 | ittell14 | B5(7)
L48 | 6 | IT1Bottom | U | 2 |  |  |  | 13 | 13 | 1 | ittell14 | B5(6)
L47 | 5 | IT1Bottom | U | 3 |  |  |  | 10 | 10 | 7 | ittell11 | B5(5)
L41 | 4 | IT1Bottom | U | 4 |  |  |  | 11 | 11 | 4 | ittell12 | B5(4)
L80 | 8 | IT1Bottom | U | 5 |  |  |  | 11 | 11 | 5 | ittell12 | B5(3)
L77 | 2 | IT1Bottom | U | 6 |  |  |  | 11 | 11 | 6 | ittell12 | B5(2)
L69 | 1 | IT1Bottom | U | 7 | [20/21] |  |  | 11 | 11 | 7 | ittell12 | B5(1)
L64 | 23 | IT1Bottom | V | 1 | [88/89];[91/92] |  |  | 13 | 13 | 4 | ittell14 | A5(7)
L60 | 24 | IT1Bottom | V | 2 |  |  |  | 13 | 13 | 2 | ittell14 | A5(6)
L44 | 26 | IT1Bottom | V | 3 |  |  |  | 13 | 13 | 3 | ittell14 | A5(5)
L57 | 25 | IT1Bottom | V | 4 |  |  |  | 12 | 12 | 0 | ittell13 | A5(4)
L61 | 27 | IT1Bottom | V | 5 |  |  |  | 12 | 12 | 1 | ittell13 | A5(3)
L53 | 28 | IT1Bottom | V | 6 |  |  |  | 12 | 12 | 2 | ittell13 | A5(2)
L120 | 29 | IT1Bottom | V | 7 |  |  |  | 12 | 12 | 3 | ittell13 | A5(1)
L54 | 12 | IT1Bottom | X1 | 1 | [121/122] |  |  | 10 | 10 | 4 | ittell11 | B0(7)
L63 | 9 | IT1Bottom | X1 | 2 |  |  |  | 10 | 10 | 5 | ittell11 | B0(6)
L42 | 13 | IT1Bottom | X1 | 3 |  |  |  | 10 | 10 | 6 | ittell11 | B0(5)
L71 | 10 | IT1Bottom | X1 | 4 |  |  |  | 11 | 11 | 0 | ittell12 | B0(4)
L45 | 14 | IT1Bottom | X1 | 5 |  |  |  | 11 | 11 | 1 | ittell12 | B0(3)
L51 | 11 | IT1Bottom | X1 | 6 |  |  |  | 11 | 11 | 2 | ittell12 | B0(2)
L55 | 15 | IT1Bottom | X1 | 7 |  |  |  | 11 | 11 | 3 | ittell12 | B0(1)
L50 | 16 | IT1Bottom | X2 | 1 |  |  |  | 13 | 13 | 5 | ittell14 | A0(7)
L62 | 17 | IT1Bottom | X2 | 2 |  |  |  | 13 | 13 | 6 | ittell14 | A0(6)
L66 | 18 | IT1Bottom | X2 | 3 |  |  |  | 13 | 13 | 7 | ittell14 | A0(5)
L70 | 19 | IT1Bottom | X2 | 4 |  |  |  | 12 | 12 | 4 | ittell13 | A0(4)
L46 | 20 | IT1Bottom | X2 | 5 |  |  |  | 12 | 12 | 5 | ittell13 | A0(3)
L79 | forgotten? | IT1Bottom | X2 | 6 |  |  |  | 12 | 12 | 6 | ittell13 | A0(2)
L59 | 22 | IT1Bottom | X2 | 7 |  |  | 236 | 12 | 12 | 7 | ittell13 | A0(1)
L246 | 59 | IT1Cryo | U | 1 |  | 20 |  | 2 | 2 | 0 | ittell3 | A5(1)
L251 | 58 | IT1Cryo | U | 2 |  |  | 64 | 2 | 2 | 1 | ittell3 | A5(2)
L249 | 57 | IT1Cryo | U | 3 |  |  |  | 3 | 3 | 3 | ittell4 | A5(3)
L93 | 56 | IT1Cryo | U | 4 |  |  |  | 0 | 0 | 4 | ittell1 | A5(4)
L247 | 55 | IT1Cryo | U | 5 |  |  |  | 0 | 0 | 5 | ittell1 | A5(5)
L101 | 54 | IT1Cryo | U | 6 | [251/252] |  | 45 | 0 | 0 | 6 | ittell1 | A5(6)
L298 | 53 | IT1Cryo | U | 7 |  |  |  | 0 | 0 | 7 | ittell1 | A5(7)
L233 | 66 | IT1Cryo | V | 1 |  |  |  | 2 | 2 | 4 | ittell3 | B5(1)
L244 | 65 | IT1Cryo | V | 2 |  |  |  | 2 | 2 | 2 | ittell3 | B5(2)
L245 | 64 | IT1Cryo | V | 3 |  |  |  | 2 | 2 | 3 | ittell3 | B5(3)
L314 | 63 | IT1Cryo | V | 4 |  |  |  | 1 | 1 | 0 | ittell2 | B5(4)
L299 | 62 | IT1Cryo | V | 5 |  |  |  | 1 | 1 | 1 | ittell2 | B5(5)
L242 | 61 | IT1Cryo | V | 6 |  |  |  | 1 | 1 | 2 | ittell2 | B5(6)
L94 | 60 | IT1Cryo | V | 7 |  |  |  | 1 | 1 | 3 | ittell2 | B5(7)
L308 | 52 | IT1Cryo | X1 | 1 |  |  |  | 3 | 3 | 0 | ittell4 | A0(1)
L276 | 51 | IT1Cryo | X1 | 2 |  |  | 383 | 3 | 3 | 1 | ittell4 | A0(2)
L304 | 132 | IT1Cryo | X1 | 3 |  |  |  | 3 | 3 | 2 | ittell4 | A0(3)
L293 | 49 | IT1Cryo | X1 | 4 |  | 371--> Was fixed!! |  | 0 | 0 | 0 | ittell1 | A0(4)
L415 | 48 | IT1Cryo | X1 | 5 |  |  |  | 0 | 0 | 1 | ittell1 | A0(5)
L231 | 47 | IT1Cryo | X1 | 6 |  |  |  | 0 | 0 | 2 | ittell1 | A0(6)
L96 | 46 | IT1Cryo | X1 | 7 |  |  |  | 0 | 0 | 3 | ittell1 | A0(7)
L291 | 73 | IT1Cryo | X2 | 1 |  |  |  | 2 | 2 | 5 | ittell3 | B0(1)
L301 | 16 | IT1Cryo | X2 | 2 |  |  |  | 2 | 2 | 6 | ittell3 | B0(2)
L106 | 71 | IT1Cryo | X2 | 3 |  |  |  | 2 | 2 | 7 | ittell3 | B0(3)
L187 | 70 | IT1Cryo | X2 | 4 |  |  |  | 1 | 1 | 4 | ittell2 | B0(4)
L222 | 69 | IT1Cryo | X2 | 5 |  |  |  | 1 | 1 | 5 | ittell2 | B0(5)
L95 | 68 | IT1Cryo | X2 | 6 | [8/9] |  |  | 1 | 1 | 6 | ittell2 | B0(6)
L234 | 67 | IT1Cryo | X2 | 7 |  |  |  | 1 | 1 | 7 | ittell2 | B0(7)
L21 | 37 | IT1Top | U | 1 |  | 375 |  | 4 | 4 | 7 | ittell5 | B5(1)
L17 | 38 | IT1Top | U | 2 | [76/77] | 225 |  | 4 | 4 | 6 | ittell5 | B5(2)
L30 | 39 | IT1Top | U | 3 |  |  |  | 4 | 4 | 5 | ittell5 | B5(3)
L18 | 40 | IT1Top | U | 4 |  |  |  | 4 | 4 | 4 | ittell5 | B5(4)
L153 | 58 | IT1Top | U | 5 |  |  | 332 | 3 | 3 | 7 | ittell4 | B5(5)
L36 | 42 | IT1Top | U | 6 |  | 145 |  | 6 | 6 | 1 | ittell7 | B5(6)
L82 | 43 | IT1Top | U | 7 |  |  |  | 6 | 6 | 0 | ittell7 | B5(7)
L22 | 44 | IT1Top | V | 1 | [353/354/355] |  |  | 5 | 5 | 3 | ittell6 | A5(1)
L133 | 59 | IT1Top | V | 2 |  |  |  | 5 | 5 | 2 | ittell6 | A5(2)
L27 | 46 | IT1Top | V | 3 | [14/15] |  | 291 | 5 | 5 | 1 | ittell6 | A5(3)
L264 | 61 | IT1Top | V | 4 |  |  |  | 5 | 5 | 0 | ittell6 | A5(4)
L34 | 48 | IT1Top | V | 5 |  |  |  | 6 | 6 | 3 | ittell7 | A5(5)
L32 | 49 | IT1Top | V | 6 | [307/308] | 115--> could be 116 |  | 6 | 6 | 2 | ittell7 | A5(6)
L38 | 50 | IT1Top | V | 7 |  |  |  | 6 | 6 | 4 | ittell7 | A5(7)
L166 | 62 | IT1Top | X1 | 1 |  |  |  | 4 | 4 | 3 | ittell5 | B0(1)
L15 | 31 | IT1Top | X1 | 2 | [20/21] |  |  | 4 | 4 | 2 | ittell5 | B0(2)
L123 | 32 | IT1Top | X1 | 3 |  |  |  | 4 | 4 | 1 | ittell5 | B0(3)
L37 | 33 | IT1Top | X1 | 4 |  |  |  | 4 | 4 | 0 | ittell5 | B0(4)
L40 | 34 | IT1Top | X1 | 5 |  |  |  | 3 | 3 | 6 | ittell4 | B0(5)
L126 | 35 | IT1Top | X1 | 6 |  |  |  | 3 | 3 | 5 | ittell4 | B0(6)
L99 | 36 | IT1Top | X1 | 7 |  |  |  | 3 | 3 | 4 | ittell4 | B0(7)
L14 | 51 | IT1Top | X2 | 1 | [365/366] |  |  | 5 | 5 | 7 | ittell6 | A0(1)
L98 | 52 | IT1Top | X2 | 2 |  | 62;78 |  | 5 | 5 | 6 | ittell6 | A0(2)
L19 | 53 | IT1Top | X2 | 3 |  |  |  | 5 | 5 | 5 | ittell6 | A0(3)
L13 | 54 | IT1Top | X2 | 4 |  |  |  | 5 | 5 | 4 | ittell6 | A0(4)
L122 | 55 | IT1Top | X2 | 5 |  |  |  | 6 | 6 | 7 | ittell7 | A0(5)
L33 | 56 | IT1Top | X2 | 6 | [210/211] |  |  | 6 | 6 | 6 | ittell7 | A0(6)
L31 | 57 | IT1Top | X2 | 7 |  |  |  | 6 | 6 | 5 | ittell7 | A0(7)
L353 | 88 | IT2Access | U | 1 |  |  |  | 7 | 39 | 4 | ittell22 | B5(7)
L104 | 89 | IT2Access | U | 2 |  |  |  | 7 | 39 | 5 | ittell22 | B5(6)
L318 | 90 | IT2Access | U | 3 |  |  |  | 7 | 39 | 6 | ittell22 | B5(5)
L319 | 91 | IT2Access | U | 4 | [30/31] |  |  | 7 | 39 | 7 | ittell22 | B5(4)
L412 | 92 | IT2Access | U | 5 | [126/127] |  |  | 10 | 42 | 1 | ittell25 | B5(3)
L349 | 93 | IT2Access | U | 6 | [98/99] |  |  | 9 | 41 | 1 | ittell24 | B5(2)
L417 | 94 | IT2Access | U | 7 | [258/257] |  |  | 9 | 41 | 3 | ittell24 | B5(1)
L338 | 81 | IT2Access | V | 1 |  |  |  | 8 | 40 | 0 | ittell23 | A5(7)
L410 | 82 | IT2Access | V | 2 |  |  |  | 8 | 40 | 1 | ittell23 | A5(6)
L339 | 17 | IT2Access | V | 3 |  |  |  | 8 | 40 | 2 | ittell23 | A5(5)
L302 | 84 | IT2Access | V | 4 | [94/95] |  | 360 | 8 | 40 | 3 | ittell23 | A5(4)
L279 | 112 | IT2Access | V | 5 |  |  |  | 9 | 41 | 0 | ittell24 | A5(3)
L278 | 107 | IT2Access | V | 6 | [36/37] |  |  | 9 | 41 | 2 | ittell24 | A5(2)
L374 | 87 | IT2Access | V | 7 |  |  |  | 9 | 41 | 6 | ittell24 | A5(1)
L403 | 95 | IT2Access | X1 | 1 |  |  |  | 7 | 39 | 0 | ittell22 | B0(7)
L402 | 96 | IT2Access | X1 | 2 |  |  |  | 7 | 39 | 1 | ittell22 | B0(6)
L399 | 97 | IT2Access | X1 | 3 | [291/292] |  |  | 7 | 39 | 2 | ittell22 | B0(5)
L350 | 98 | IT2Access | X1 | 4 |  |  |  | 7 | 39 | 3 | ittell22 | B0(4)
L395 | 99 | IT2Access | X1 | 5 |  |  |  | 10 | 42 | 0 | ittell25 | B0(3)
L418 | 100 | IT2Access | X1 | 6 |  | 353;354 |  | 10 | 42 | 2 | ittell25 | B0(2)
L392 | 101 | IT2Access | X1 | 7 |  |  |  | 10 | 42 | 3 | ittell25 | B0(1)
L352 | 74 | IT2Access | X2 | 1 |  |  |  | 8 | 40 | 4 | ittell23 | A0(7)
L394 | 75 | IT2Access | X2 | 2 |  |  |  | 8 | 40 | 5 | ittell23 | A0(6)
L366 | 76 | IT2Access | X2 | 3 |  |  |  | 8 | 40 | 6 | ittell23 | A0(5)
L396 | 77 | IT2Access | X2 | 4 |  |  | 159 | 8 | 40 | 7 | ittell23 | A0(4)
L269 | 78 | IT2Access | X2 | 5 |  |  |  | 9 | 41 | 4 | ittell24 | A0(3)
L272 | 79 | IT2Access | X2 | 6 | [331/332] | 63-->had been fixed |  | 9 | 41 | 5 | ittell24 | A0(2)
L421 | 80 | IT2Access | X2 | 7 | [331/332] |  |  | 9 | 41 | 7 | ittell24 | A0(1)
L165 | 140 | IT2Bottom | U | 1 |  |  |  | 13 | 45 | 0 | ittell28 | B5(7)
L400 | 139 | IT2Bottom | U | 2 |  |  |  | 13 | 45 | 1 | ittell28 | B5(6)
L124 | 138 | IT2Bottom | U | 3 |  |  |  | 10 | 42 | 7 | ittell25 | B5(5)
L129 | 137 | IT2Bottom | U | 4 |  |  |  | 11 | 43 | 4 | ittell26 | B5(4)
L330 | 136 | IT2Bottom | U | 5 | [113/114] |  |  | 11 | 43 | 5 | ittell26 | B5(3)
L326 | 135 | IT2Bottom | U | 6 |  |  |  | 11 | 43 | 6 | ittell26 | B5(2)
L254 | 134 | IT2Bottom | U | 7 |  | 137 |  | 11 | 43 | 7 | ittell26 | B5(1)
L267 | 133 | IT2Bottom | V | 1 |  |  |  | 13 | 45 | 4 | ittell28 | A5(7)
L136 | 150 | IT2Bottom | V | 2 |  |  |  | 13 | 45 | 2 | ittell28 | A5(6)
L125 | 131 | IT2Bottom | V | 3 |  |  |  | 13 | 45 | 3 | ittell28 | A5(5)
L65 | 130 | IT2Bottom | V | 4 |  | 293 Open or Pinhole? | 293 | 12 | 44 | 0 | ittell27 | A5(4)
L89 | 129 | IT2Bottom | V | 5 |  |  |  | 12 | 44 | 1 | ittell27 | A5(3)
L81 | 128 | IT2Bottom | V | 6 |  |  |  | 12 | 44 | 2 | ittell27 | A5(2)
L86 | 127 | IT2Bottom | V | 7 |  |  |  | 12 | 44 | 3 | ittell27 | A5(1)
L151 | 141 | IT2Bottom | X1 | 1 |  |  |  | 10 | 42 | 4 | ittell25 | B0(7)
L331 | 146 | IT2Bottom | X1 | 2 |  | 327-->had been fixed |  | 10 | 42 | 5 | ittell25 | B0(6)
L138 | 143 | IT2Bottom | X1 | 3 |  |  |  | 10 | 42 | 6 | ittell25 | B0(5)
L386 | 144 | IT2Bottom | X1 | 4 |  |  |  | 11 | 43 | 0 | ittell26 | B0(4)
L29 | 47 | IT2Bottom | X1 | 5 |  |  |  | 11 | 43 | 1 | ittell26 | B0(3)
L143 | 142 | IT2Bottom | X1 | 6 |  |  |  | 11 | 43 | 2 | ittell26 | B0(2)
L139 | 147 | IT2Bottom | X1 | 7 |  |  |  | 11 | 43 | 3 | ittell26 | B0(1)
L43 | 126 | IT2Bottom | X2 | 1 |  | 163 |  | 13 | 45 | 5 | ittell28 | A0(7)
L174 | 125 | IT2Bottom | X2 | 2 |  |  |  | 13 | 45 | 6 | ittell28 | A0(6)
L134 | 124 | IT2Bottom | X2 | 3 |  |  |  | 13 | 45 | 7 | ittell28 | A0(5)
L154 | 123 | IT2Bottom | X2 | 4 |  |  |  | 12 | 44 | 4 | ittell27 | A0(4)
L130 | 122 | IT2Bottom | X2 | 5 |  |  |  | 12 | 44 | 5 | ittell27 | A0(3)
L167 | 121 | IT2Bottom | X2 | 6 |  |  |  | 12 | 44 | 6 | ittell27 | A0(2)
L334 | 120 | IT2Bottom | X2 | 7 | [5/6] |  |  | 12 | 44 | 7 | ittell27 | A0(1)
L337 | 122 | IT2Cryo | U | 1 |  |  |  | 2 | 34 | 0 | ittell17 | A5(1)
L325 | 123 | IT2Cryo | U | 2 |  |  |  | 2 | 34 | 1 | ittell17 | A5(2)
L282 | 106 | IT2Cryo | U | 3 |  |  |  | 3 | 35 | 3 | ittell18 | A5(3)
L347 | 125 | IT2Cryo | U | 4 |  |  |  | 0 | 32 | 4 | ittell15 | A5(4)
L340 | 126 | IT2Cryo | U | 5 |  |  |  | 0 | 32 | 5 | ittell15 | A5(5)
L341 | 127 | IT2Cryo | U | 6 |  |  |  | 0 | 32 | 6 | ittell15 | A5(6)
L342 | 102 | IT2Cryo | U | 7 |  |  |  | 0 | 32 | 7 | ittell15 | A5(7)
L307 | 108 | IT2Cryo | V | 1 |  |  |  | 2 | 34 | 4 | ittell17 | B5(1)
L283 | 109 | IT2Cryo | V | 2 |  |  |  | 2 | 34 | 2 | ittell17 | B5(2)
L306 | 110 | IT2Cryo | V | 3 |  |  |  | 2 | 34 | 3 | ittell17 | B5(3)
L275 | 111 | IT2Cryo | V | 4 |  |  |  | 1 | 33 | 0 | ittell16 | B5(4)
L303 | 104 | IT2Cryo | V | 5 |  |  |  | 1 | 33 | 1 | ittell16 | B5(5)
L336 | 113 | IT2Cryo | V | 6 |  |  |  | 1 | 33 | 2 | ittell16 | B5(6)
L343 | 114 | IT2Cryo | V | 7 |  |  |  | 1 | 33 | 3 | ittell16 | B5(7)
L277 | 129 | IT2Cryo | X1 | 1 |  |  |  | 3 | 35 | 0 | ittell18 | A0(1)
L312 | 130 | IT2Cryo | X1 | 2 |  |  |  | 3 | 35 | 1 | ittell18 | A0(2)
L288 | 131 | IT2Cryo | X1 | 3 |  |  |  | 3 | 35 | 2 | ittell18 | A0(3)
L281 | 105 | IT2Cryo | X1 | 4 |  | 149 |  | 0 | 32 | 0 | ittell15 | A0(4)
L286 | 133 | IT2Cryo | X1 | 5 |  |  |  | 0 | 32 | 1 | ittell15 | A0(5)
L287 | 134 | IT2Cryo | X1 | 6 |  |  |  | 0 | 32 | 2 | ittell15 | A0(6)
L268 | 135 | IT2Cryo | X1 | 7 |  |  |  | 0 | 32 | 3 | ittell15 | A0(7)
L309 | 115 | IT2Cryo | X2 | 1 |  |  |  | 2 | 34 | 5 | ittell17 | B0(1)
L284 | 116 | IT2Cryo | X2 | 2 |  |  |  | 2 | 34 | 6 | ittell17 | B0(2)
L230 | 160 | IT2Cryo | X2 | 3 |  |  |  | 2 | 34 | 7 | ittell17 | B0(3)
L393 | 103 | IT2Cryo | X2 | 4 |  |  |  | 1 | 33 | 4 | ittell16 | B0(4)
L362 | 119 | IT2Cryo | X2 | 5 |  |  |  | 1 | 33 | 5 | ittell16 | B0(5)
L361 | 120 | IT2Cryo | X2 | 6 |  |  |  | 1 | 33 | 6 | ittell16 | B0(6)
L365 | 121 | IT2Cryo | X2 | 7 |  |  |  | 1 | 33 | 7 | ittell16 | B0(7)
L260 | 69 | IT2Top | U | 1 |  |  | 280 | 4 | 36 | 7 | ittell19 | B5(1)
L128 | 68 | IT2Top | U | 2 |  | 16 |  | 4 | 36 | 6 | ittell19 | B5(2)
L160 | 67 | IT2Top | U | 3 |  |  |  | 4 | 36 | 5 | ittell19 | B5(3)
L384 | 75 | IT2Top | U | 4 |  |  |  | 4 | 36 | 4 | ittell19 | B5(4)
L97 | 65 | IT2Top | U | 5 |  |  |  | 3 | 35 | 7 | ittell18 | B5(5)
L164 | 64 | IT2Top | U | 6 |  |  |  | 6 | 38 | 1 | ittell21 | B5(6)
L257 | 63 | IT2Top | U | 7 |  |  | 20;233 | 6 | 38 | 0 | ittell21 | B5(7)
L172 | 77 | IT2Top | V | 1 |  |  |  | 5 | 37 | 3 | ittell20 | A5(1)
L144 | 78 | IT2Top | V | 2 |  |  |  | 5 | 37 | 2 | ittell20 | A5(2)
L145 | 79 | IT2Top | V | 3 |  |  |  | 5 | 37 | 1 | ittell20 | A5(3)
L377 | 91 | IT2Top | V | 4 |  |  |  | 5 | 37 | 0 | ittell20 | A5(4)
L121 | 81 | IT2Top | V | 5 | [285/286] |  |  | 6 | 38 | 3 | ittell21 | A5(5)
L39 | 82 | IT2Top | V | 6 | [288/289] |  |  | 6 | 38 | 2 | ittell21 | A5(6)
L265 | 83 | IT2Top | V | 7 |  | 84 |  | 6 | 38 | 4 | ittell21 | A5(7)
L146 | 76 | IT2Top | X1 | 1 |  |  |  | 4 | 36 | 3 | ittell19 | B0(1)
L24 | 45 | IT2Top | X1 | 2 |  |  |  | 4 | 36 | 2 | ittell19 | B0(2)
L373 | 74 | IT2Top | X1 | 3 |  |  |  | 4 | 36 | 1 | ittell19 | B0(3)
L170 | 73 | IT2Top | X1 | 4 |  |  |  | 4 | 36 | 0 | ittell19 | B0(4)
L178 | 72 | IT2Top | X1 | 5 |  |  |  | 3 | 35 | 6 | ittell18 | B0(5)
L158 | 71 | IT2Top | X1 | 6 |  |  |  | 3 | 35 | 5 | ittell18 | B0(6)
L364 | 88 | IT2Top | X1 | 7 | [300/301];[308/309] |  | 309 | 3 | 35 | 4 | ittell18 | B0(7)
L175 | 117 | IT2Top | X2 | 1 |  | 1 |  | 5 | 37 | 7 | ittell20 | A0(1)
L179 | 85 | IT2Top | X2 | 2 |  |  |  | 5 | 37 | 6 | ittell20 | A0(2)
L159 | 86 | IT2Top | X2 | 3 |  |  |  | 5 | 37 | 5 | ittell20 | A0(3)
L171 | 87 | IT2Top | X2 | 4 |  |  |  | 5 | 37 | 4 | ittell20 | A0(4)
L78 | 21 | IT2Top | X2 | 5 |  |  |  | 6 | 38 | 7 | ittell21 | A0(5)
L333 | 89 | IT2Top | X2 | 6 |  |  |  | 6 | 38 | 6 | ittell21 | A0(6)
L328 | 90 | IT2Top | X2 | 7 |  |  |  | 6 | 38 | 5 | ittell21 | A0(7)
L205 | 151 | IT3Access | U | 1 |  |  |  | 7 | 71 | 4 | ittell36 | B5(7)
L208 | 152 | IT3Access | U | 2 |  |  |  | 7 | 71 | 5 | ittell36 | B5(6)
L220 | 153 | IT3Access | U | 3 |  |  |  | 7 | 71 | 6 | ittell36 | B5(5)
L221 | 154 | IT3Access | U | 4 |  |  |  | 7 | 71 | 7 | ittell36 | B5(4)
L225 | 155 | IT3Access | U | 5 |  |  |  | 10 | 74 | 1 | ittell39 | B5(3)
L217 | 156 | IT3Access | U | 6 |  | 146 |  | 9 | 73 | 1 | ittell38 | B5(2)
L216 | 157 | IT3Access | U | 7 |  |  |  | 9 | 73 | 3 | ittell38 | B5(1)
L200 | 144 | IT3Access | V | 1 |  |  |  | 8 | 72 | 0 | ittell37 | A5(7)
L204 | 145 | IT3Access | V | 2 |  |  |  | 8 | 72 | 1 | ittell37 | A5(6)
L209 | 146 | IT3Access | V | 3 |  |  |  | 8 | 72 | 2 | ittell37 | A5(5)
L212 | 147 | IT3Access | V | 4 |  |  |  | 8 | 72 | 3 | ittell37 | A5(4)
L239 | 148 | IT3Access | V | 5 | [106/107] |  |  | 9 | 73 | 0 | ittell38 | A5(3)
L238 | 149 | IT3Access | V | 6 | [107/108] |  |  | 9 | 73 | 2 | ittell38 | A5(2)
L213 | 150 | IT3Access | V | 7 |  |  |  | 9 | 73 | 6 | ittell38 | A5(1)
L206 | 158 | IT3Access | X1 | 1 |  |  |  | 7 | 71 | 0 | ittell36 | B0(7)
L218 | 159 | IT3Access | X1 | 2 |  |  |  | 7 | 71 | 1 | ittell36 | B0(6)
L285 | 136 | IT3Access | X1 | 3 |  |  |  | 7 | 71 | 2 | ittell36 | B0(5)
L207 | 161 | IT3Access | X1 | 4 | [16/17] |  |  | 7 | 71 | 3 | ittell36 | B0(4)
L227 | 162 | IT3Access | X1 | 5 |  |  |  | 10 | 74 | 0 | ittell39 | B0(3)
L219 | 163 | IT3Access | X1 | 6 | [108/109] Not seen in pit data? |  |  | 10 | 74 | 2 | ittell39 | B0(2)
L226 | 164 | IT3Access | X1 | 7 | [361/362] |  |  | 10 | 74 | 3 | ittell39 | B0(1)
L198 | 137 | IT3Access | X2 | 1 |  |  |  | 8 | 72 | 4 | ittell37 | A0(7)
L199 | 138 | IT3Access | X2 | 2 |  |  |  | 8 | 72 | 5 | ittell37 | A0(6)
L202 | 139 | IT3Access | X2 | 3 |  |  |  | 8 | 72 | 6 | ittell37 | A0(5)
L203 | 140 | IT3Access | X2 | 4 |  |  |  | 8 | 72 | 7 | ittell37 | A0(4)
L211 | 141 | IT3Access | X2 | 5 |  |  |  | 9 | 73 | 4 | ittell38 | A0(3)
L215 | 142 | IT3Access | X2 | 6 |  |  |  | 9 | 73 | 5 | ittell38 | A0(2)
L214 | 143 | IT3Access | X2 | 7 |  |  | 121 | 9 | 73 | 7 | ittell38 | A0(1)
L388 | 168 | IT3Bottom | U | 1 |  |  |  | 13 | 77 | 0 | ittell42 | B5(7)
L258 | 167 | IT3Bottom | U | 2 |  |  |  | 13 | 77 | 1 | ittell42 | B5(6)
L380 | 166 | IT3Bottom | U | 3 |  | 92 | 92 | 10 | 74 | 7 | ittell39 | B5(5)
L148 | 176 | IT3Bottom | U | 4 |  |  |  | 11 | 75 | 4 | ittell40 | B5(4)
L156 | 164 | IT3Bottom | U | 5 |  |  |  | 11 | 75 | 5 | ittell40 | B5(3)
L261 | 163 | IT3Bottom | U | 6 |  |  |  | 11 | 75 | 6 | ittell40 | B5(2)
L169 | 162 | IT3Bottom | U | 7 |  |  |  | 11 | 75 | 7 | ittell40 | B5(1)
L408 | 161 | IT3Bottom | V | 1 |  |  |  | 13 | 77 | 4 | ittell42 | A5(7)
L381 | 160 | IT3Bottom | V | 2 |  |  |  | 13 | 77 | 2 | ittell42 | A5(6)
L259 | 159 | IT3Bottom | V | 3 |  |  |  | 13 | 77 | 3 | ittell42 | A5(5)
L176 | 158 | IT3Bottom | V | 4 |  |  |  | 12 | 76 | 0 | ittell41 | A5(4)
L173 | 157 | IT3Bottom | V | 5 |  |  |  | 12 | 76 | 1 | ittell41 | A5(3)
L132 | 156 | IT3Bottom | V | 6 |  |  |  | 12 | 76 | 2 | ittell41 | A5(2)
L168 | 155 | IT3Bottom | V | 7 |  |  |  | 12 | 76 | 3 | ittell41 | A5(1)
L387 | 175 | IT3Bottom | X1 | 1 |  |  |  | 10 | 74 | 4 | ittell39 | B0(7)
L35 | 174 | IT3Bottom | X1 | 2 |  | 174 |  | 10 | 74 | 5 | ittell39 | B0(6)
L356 | 173 | IT3Bottom | X1 | 3 |  |  |  | 10 | 74 | 6 | ittell39 | B0(5)
L6 | 172 | IT3Bottom | X1 | 4 | [151/152];[353/354] |  |  | 11 | 75 | 0 | ittell40 | B0(4)
L359 | 171 | IT3Bottom | X1 | 5 |  |  |  | 11 | 75 | 1 | ittell40 | B0(3)
L358 | 170 | IT3Bottom | X1 | 6 |  |  |  | 11 | 75 | 2 | ittell40 | B0(2)
L357 | 169 | IT3Bottom | X1 | 7 |  |  |  | 11 | 75 | 3 | ittell40 | B0(1)
L127 | 154 | IT3Bottom | X2 | 1 |  |  |  | 13 | 77 | 5 | ittell42 | A0(7)
L131 | 153 | IT3Bottom | X2 | 2 |  |  |  | 13 | 77 | 6 | ittell42 | A0(6)
L335 | 152 | IT3Bottom | X2 | 3 |  |  |  | 13 | 77 | 7 | ittell42 | A0(5)
L3 | 151 | IT3Bottom | X2 | 4 | [70/71] |  |  | 12 | 76 | 4 | ittell41 | A0(4)
L87 | 41 | IT3Bottom | X2 | 5 |  |  |  | 12 | 76 | 5 | ittell41 | A0(3)
L363 | 149 | IT3Bottom | X2 | 6 |  |  |  | 12 | 76 | 6 | ittell41 | A0(2)
L360 | 148 | IT3Bottom | X2 | 7 |  |  |  | 12 | 76 | 7 | ittell41 | A0(1)
L323 | 15 | IT3Cryo | U | 1 |  |  |  | 2 | 66 | 0 | ittell31 | A5(1)
L224 | 30 | IT3Cryo | U | 2 |  |  |  | 2 | 66 | 1 | ittell31 | A5(2)
L274 | 29 | IT3Cryo | U | 3 |  |  |  | 3 | 67 | 3 | ittell32 | A5(3)
L229 | 28 | IT3Cryo | U | 4 | [45/46/47] |  |  | 0 | 64 | 4 | ittell29 | A5(4)
L236 | 27 | IT3Cryo | U | 5 |  |  | 292 | 0 | 64 | 5 | ittell29 | A5(5)
L420 | 26 | IT3Cryo | U | 6 |  |  |  | 0 | 64 | 6 | ittell29 | A5(6)
L180 | 25 | IT3Cryo | U | 7 |  |  |  | 0 | 64 | 7 | ittell29 | A5(7)
L181 | 38 | IT3Cryo | V | 1 |  | 275 |  | 2 | 66 | 4 | ittell31 | B5(1)
L201 | 37 | IT3Cryo | V | 2 | [95/97] |  |  | 2 | 66 | 2 | ittell31 | B5(2)
L100 | 36 | IT3Cryo | V | 3 |  |  |  | 2 | 66 | 3 | ittell31 | B5(3)
L237 | 35 | IT3Cryo | V | 4 |  |  |  | 1 | 65 | 0 | ittell30 | B5(4)
L241 | 34 | IT3Cryo | V | 5 |  |  |  | 1 | 65 | 1 | ittell30 | B5(5)
L243 | 33 | IT3Cryo | V | 6 |  |  |  | 1 | 65 | 2 | ittell30 | B5(6)
L232 | 32 | IT3Cryo | V | 7 |  |  |  | 1 | 65 | 3 | ittell30 | B5(7)
L273 | 24 | IT3Cryo | X1 | 1 |  |  |  | 3 | 67 | 0 | ittell32 | A0(1)
L348 | 23 | IT3Cryo | X1 | 2 |  | 149 |  | 3 | 67 | 1 | ittell32 | A0(2)
L344 | 22 | IT3Cryo | X1 | 3 |  |  |  | 3 | 67 | 2 | ittell32 | A0(3)
L295 | 21 | IT3Cryo | X1 | 4 |  |  |  | 0 | 64 | 0 | ittell29 | A0(4)
L297 | 20 | IT3Cryo | X1 | 5 |  |  |  | 0 | 64 | 1 | ittell29 | A0(5)
L103 | 19 | IT3Cryo | X1 | 6 |  |  |  | 0 | 64 | 2 | ittell29 | A0(6)
L409 | 18 | IT3Cryo | X1 | 7 |  |  |  | 0 | 64 | 3 | ittell29 | A0(7)
L398 | 45 | IT3Cryo | X2 | 1 |  |  |  | 2 | 66 | 5 | ittell31 | B0(1)
L235 | 44 | IT3Cryo | X2 | 2 |  |  |  | 2 | 66 | 6 | ittell31 | B0(2)
L296 | 43 | IT3Cryo | X2 | 3 |  |  | 248 | 2 | 66 | 7 | ittell31 | B0(3)
L294 | 42 | IT3Cryo | X2 | 4 |  |  |  | 1 | 65 | 4 | ittell30 | B0(4)
L397 | 41 | IT3Cryo | X2 | 5 | [107/108] |  |  | 1 | 65 | 5 | ittell30 | B0(5)
L305 | 118 | IT3Cryo | X2 | 6 |  |  |  | 1 | 65 | 6 | ittell30 | B0(6)
L271 | 39 | IT3Cryo | X2 | 7 |  |  |  | 1 | 65 | 7 | ittell30 | B0(7)
L263 | 106 | IT3Top | U | 1 |  |  |  | 4 | 68 | 7 | ittell33 | B5(1)
L152 | 107 | IT3Top | U | 2 |  |  |  | 4 | 68 | 6 | ittell33 | B5(2)
L401 | 108 | IT3Top | U | 3 |  |  |  | 4 | 68 | 5 | ittell33 | B5(3)
L177 | 109 | IT3Top | U | 4 |  |  |  | 4 | 68 | 4 | ittell33 | B5(4)
L141 | 110 | IT3Top | U | 5 |  |  |  | 3 | 67 | 7 | ittell32 | B5(5)
L407 | 111 | IT3Top | U | 6 |  |  |  | 6 | 70 | 1 | ittell35 | B5(6)
L149 | 112 | IT3Top | U | 7 |  |  |  | 6 | 70 | 0 | ittell35 | B5(7)
L90 | 99 | IT3Top | V | 1 |  |  |  | 5 | 69 | 3 | ittell34 | A5(1)
L72 | 100 | IT3Top | V | 2 |  |  |  | 5 | 69 | 2 | ittell34 | A5(2)
L68 | 101 | IT3Top | V | 3 |  |  |  | 5 | 69 | 1 | ittell34 | A5(3)
L76 | 102 | IT3Top | V | 4 |  |  |  | 5 | 69 | 0 | ittell34 | A5(4)
L405 | 103 | IT3Top | V | 5 |  |  |  | 6 | 70 | 3 | ittell35 | A5(5)
L137 | 177 | IT3Top | V | 6 |  |  |  | 6 | 70 | 2 | ittell35 | A5(6)
L157 | 105 | IT3Top | V | 7 |  |  |  | 6 | 70 | 4 | ittell35 | A5(7)
L150 | 178 | IT3Top | X1 | 1 |  |  |  | 4 | 68 | 3 | ittell33 | B0(1)
L147 | 114 | IT3Top | X1 | 2 |  |  |  | 4 | 68 | 2 | ittell33 | B0(2)
L329 | 115 | IT3Top | X1 | 3 |  |  |  | 4 | 68 | 1 | ittell33 | B0(3)
L354 | 116 | IT3Top | X1 | 4 |  |  |  | 4 | 68 | 0 | ittell33 | B0(4)
L162 | 117 | IT3Top | X1 | 5 | [260/261] |  |  | 3 | 67 | 6 | ittell32 | B0(5)
L368 | 118 | IT3Top | X1 | 6 |  |  |  | 3 | 67 | 5 | ittell32 | B0(6)
L382 | 119 | IT3Top | X1 | 7 |  |  |  | 3 | 67 | 4 | ittell32 | B0(7)
L67 | 92 | IT3Top | X2 | 1 | [9/10] |  |  | 5 | 69 | 7 | ittell34 | A0(1)
L75 | 93 | IT3Top | X2 | 2 |  |  |  | 5 | 69 | 6 | ittell34 | A0(2)
L88 | 94 | IT3Top | X2 | 3 |  |  |  | 5 | 69 | 5 | ittell34 | A0(3)
L84 | 95 | IT3Top | X2 | 4 |  |  |  | 5 | 69 | 4 | ittell34 | A0(4)
L371 | 96 | IT3Top | X2 | 5 |  |  |  | 6 | 70 | 7 | ittell35 | A0(5)
L376 | 97 | IT3Top | X2 | 6 |  | 130 Not seen in pit data |  | 6 | 70 | 6 | ittell35 | A0(6)
L327 | 98 | IT3Top | X2 | 7 |  |  |  | 6 | 70 | 5 | ittell35 | A0(7)